## Slide 1: Electrical Safety Testing
Planned preventative maintenance

## Slide 2
Overview
Scope of Presentation
Background
Standards & Classifications
Testing & Considerations
Summary & Closing Comments
Overview
04

## Slide 3
Scope of Presentation
05
Scope of Presentation

## Slide 4
Discuss various test criteria, classifications, terminology, and standards.
Not a directive on what you should do in your facility
Scope of Presentation
06

## Slide 5: Medical Products - Life Cycle
R&D
Production
Acceptance
PPM & repair
Decom’ing
Realm of the Biomed
Background
07

## Slide 6: Safety testing - Why?
Objective to test and confirm the medical equipment is electrically safe to be used in a healthcare environment.
The most frequent and common testing a biomed does
Background
08

## Slide 7: Safety testing - Why?
1970s a great deal of attention was paid to “micro shock hazard”
Studies showed that ventricular fibrillation could be introduced to the patient
History
Background
09

## Slide 8: Safety testing - Why?
1971 article by Ralph Nader in Lady’s Home Journal
Background
10

## Slide 9: Preventative Maintenance – PM
Objective is to prevent or find potential faults
Visual Inspection
Testing means of protection
  Line cord
  Ground continuity
  Insulation
  Leakage currents
  Performance testing
Data capture and management
Background
11

## Slide 10: Electrical current
Electrical Current;
  Requires a closed circuit
  Follows the path of least resistance (much like water)
  Electrical (functional) currents form an integral part during monitoring, treatment and diagnosis of patients
But can cause injury or even result in death as a result of muscle stimulation, tissue burns or trigger other events (e.g., ignition of pure oxygen)
Risk of (non-functional) electrical leakage is relatively small but always present.
Life saver in defibrillators (functional current)
Background
12

## Slide 11: Electrical current on human body
Impact of current on the human body, adapted from research by Howard M. Hochberg, 1971
Background
13

## Slide 12
Macroshock – an electrical current passing through the body when contact is made between skin and an electrical source (non-invasive)
Physiological Effect | Gender | DC | 60Hz AC | 10 kHz AC
Slight sensation | Men | 1 mA | 0.4 mA | 7 mA
 | Women | 0.6 mA | 0.3 mA | 5 mA
Threshold of perception | Men | 5.2 mA | 1.1 mA | 12 mA
 | Women | 3.5 mA | 0.7 mA | 8 mA
Pain, voluntary muscle control "Let-go" | Men | 62 mA | 9 mA | 55 mA
 | Women | 41 mA | 6 mA | 37 mA
Pain, involuntary muscle control | Men | 76 mA | 16 mA | 75 mA
 | Women | 51 mA | 10.5 mA | 50 mA
Severe pain. Difficulty breathing 99.5% muscle control lost | Men | 90 mA | 23 mA | 94 mA
 | Women | 60 mA | 15 mA | 63 mA
Ventricular fibrillation (3 seconds) | Men | 500 mA | 100 mA | 
 | Women | 500 mA | 100 mA | 
http://highvoltageconnection.com/articles/ElectricShockQuestions.htm
14

## Slide 13
(C.F. Dalziel, “Electrical Shock”, Advances in Biomedical Engineering, edited by J. H. U. Brown and J. F. Dickson III, 1973, 3, 223-248)
The average resistance of the human body can be ~ 1 - 2 kΩ
There is no physiological difference between “Let-go” currents at commercial frequencies of 50 to 60Hz
The frequency response between 5 to 1000 Hz is relatively flat
As the current frequency increases, the body becomes less stimulated
15

## Slide 14
Microshock – current following directly through the heart (invasive)
Shock risk from:
  Intracardiac connections – pacemakers or saline catheters
Current (μA) | Physiological Effect
20 - 100 | Ventricular Fibrillation, death
http://www.ccanesthesiareview.com/QF/Electrical_Shock.pdf
16

## Slide 15: Leakage Current
“Current that is not functional”
  IEC 60601-1:2006
Unavoidable and as a result of capacitive and resistive dielectrics
Flowing from high to low potential ( to ground or across the body arm – to arm)
Non-hazardous if controlled through IEC 60601 design criteria
Potentially hazardous if outside the stated standard limits
Easily overlooked and subject to incorrect testing
Background
17

## Slide 16
Standards & Classifications
18
Standards & Classifications

## Slide 17: Standards
IEC60601 (ES60601)
IEC62353
NFPA 99
IEC 61010 (UL 61010)
Tests
Standards & Classifications
19
IEC 61010 - Safety requirements for electrical equipment for measurement, control, and laboratory use NOT MEDICAL EQUIPMENT
  The definition of a Medical device is: “any instrument, apparatus, implement, machine, appliance, implant, reagent for in vitro use, software, material or other similar, intended by the manufacturer to be used, alone or in combination, for human beings, for one or more of the specific medical purposes to diagnose, prevent, treat and alleviate illness/injury, and support life, ” (http://www.who.int/medical_devices/full_deffinition/en/)

## Slide 18: Local adaptations of “field” standards
PM
IEC62353
VDE 0751 / 62353(Germany / Austria)
MDA DB 2006 (UK)
IPEM Report 97 (UK)
NFPA 99 (USA)
AS/NZ 3551 (Australia)
Malaysia MS 62353
Standards & Classifications
20

## Slide 19: Classifications IEC 60601-1
Class I
Class I
Equipment protection against electric shock is achieved through means of connecting exposed conductive parts to the protective earth in the fixed wiring of the installation or through the grounded power cord of portable equipment..
Standards & Classifications
21

## Slide 20: Classifications
Class II
Class II
Also referred to as double insulated. Equipment protection against electric shock is achieved by additional protection to basic insulation through means of supplementary insulation, there being no provision for the connection of exposed metalwork of the equipment to a protective conductor and no reliance upon precautions to be taken in the fixed wiring of the installation.
Classifications are not to be confused with FDA equipment classifications.
Standards & Classifications
22

## Slide 21: ME Class (input protection) IEC 60601-1
ME class 1 (Protection relying on fault currents to ground (earth)
ME class 2 (Double Insulated, protection relying on additional insulation)
Standards & Classifications
23

## Slide 22: ME Type (output protection)
Type B (body)- Non-invasive. Earth Referenced
Type BF – Non Invasive. Patient Circuit Floating type.
Type CF – Cardiac. Patient Circuit Floating type.
Standards & Classifications
24

## Slide 23: Examples of Applied Parts – Patient Touch Points – Patient Leads
ECG module (CF)
  Patient connections are the leads 3-10
Defib paddle module (BF)
  Patient connections are the paddles
SPO2 (BF)
Infusion Device (CF)
Hospital Bed (B)
Standards & Classifications
25

## Slide 24: IEC 60601-1
Design and Type-Test standard (Mandatory)
Electrical and Mechanical Safety and Compliance of Medical Electronic Equipment
Ensure Safety of Patient / User and Environment
First published in 1976/77
Standards & Classifications
26

## Slide 25: Single fault conditions in 60601
Open earth
Open neutral
Mains on applied parts
  Applied parts = patient connections
Mains on signal input / signal output
Standards & Classifications
27

## Slide 26: Applied Part testing in 60601
Applied Part - Part of Medical Electrical Equipment that in normal use necessarily comes into physical contact with the patient
All tested to ground individually
All tested lead-to-lead individually
Under normal and single fault conditions
Standards & Classifications
28

## Slide 27: 62353
IEC standard for field testing
  Direct
  Alternative
  Differential
Not a 60601 replacement, but a test protocol for field located equipment. (Already qualified to 60601 and now tested to 62353)
It seeks to provide a uniform and unambiguous means of assessing the safety of medical equipment, while maintaining the relation to IEC 60601-1
Standards & Classifications
29

## Slide 28: NFPA99
NFPA99-2018 is the latest
Common in US hospitals
Dictated by law in many states
Standards & Classifications
30

## Slide 29: Classifications of Areas
NFPA 99 requires a healthcare facility’s governing body to designate patient care rooms within the facility
NFPA 1.3.4.1
It will develop medical device risk-assessment maintenance criteria
Standards & Classifications
31

## Slide 30: Classifications of Areas
NFPA 99 requires a healthcare facility’s governing body to designate patient care rooms within the facility
NFPA 1.3.4.1
It will develop medical device risk-assessment maintenance criteria
Standards & Classifications
32

## Slide 31: Category 1
A room in which the failure of equipment or a system is likely to cause major injury or death of patients or caregivers.
  ICU
  CCU
  Cath Labs
  OR
  ER
NFPA 99
Standards & Classifications
33

## Slide 32: Category 2
A room in which failure of equipment or a system is likely to cause minor injury to patients or caregivers.
  Inpatient bedrooms
  Dialysis Rooms
  Procedural Rooms
NFPA 99
Standards & Classifications
34

## Slide 33: Category 3
A room in which failure of equipment or a system is NOT likely to cause injury to patients or caregivers.
  Exam Rooms
  Dental Office
NFPA 99
Standards & Classifications
35

## Slide 34: Category 4
A room in which failure of equipment or a system is not likely to have a physical impact on patients or caregivers.
  Labs
  Waiting Rooms
  Morgues
NFPA 99
Standards & Classifications
36

## Slide 35: Classification
Class II - Double insulated appliances shall be permitted to have 2 conductor cords and shall be rated as Class II devices
Class II equipment is permitted to have two prong cords in the patient care vicinity
NFPA 99
Standards & Classifications
37

## Slide 36: Fixed Equipment
It is equipment permanently wired to a certain location.
  In patient care vicinity should tested before installation.
  The leakage current should not exceed 10mA
NFPA 99
Standards & Classifications
38

## Slide 37: Portable Medical Equipment
Portable electrically powered equipment that can be moved from location to location if necessary.
Standards & Classifications
39

## Slide 38
Testing & considerations
40
Testing & Considerations

## Slide 39: Ground Resistance
The power cord ground resistance is measured by connecting one lead of the plug grounding pin and the other to an exposed conductive surface
Grounding pin on the equipment is the preferred measuring point
Portable Equipment
Testing & Considerations
41

## Slide 40: Ground Resistance
Set up
Testing & Considerations
42

## Slide 41: NFPA99
10.3.2* Resistance.
10.3.2.1 For appliances that are used in the patient care vicinity, the resistance between the appliance chassis, or any exposed conductive surface of the appliance, and the ground pin of the attachment plug shall be less than 0.50 ohm under the following conditions:
  (1) The cord shall be flexed at its connection to the attachment plug or connector.
  (2) The cord shall be flexed at its connection to the strain relief on the chassis.
10.3.2.2 The requirement of 10.3.2.1 shall not apply to accessible metal parts that achieve separation from main parts by double insulation or metallic screening or that are unlikely to become energized (e.g., escutcheons or nameplates, small screws).
Testing & Considerations
43

## Slide 42: Quick view table
Ground resistance
Zero readings are suspect.
Testing & Considerations
44

## Slide 43: Touch Current or Leakage Current
New term for “leakage current” is “touch current”.
Reason? To remain consistent with International Standards (IEC).
Portable Equipment
Testing & Considerations
45

## Slide 44: Touch Current
It is current flowing from the enclosure that is accessible to the operator or patient.
Must be tested with a Safety Analyzer from a conductive surface to the ground
Typically by connecting the meter to the ground of an electrical receptacle in good conditions
Portable Equipment
Testing & Considerations
46

## Slide 45: Touch Current
Measurements are made under four conditions
  Ground closed and power off
  Ground closed and power on
  Ground open and power off
  Ground open and power on
Portable Equipment
Testing & Considerations
47

## Slide 46: NFPA99
10.3.4 Leakage Current — Fixed Equipment.
10.3.4.1 Permanently wired appliances in the patient care vicinity shall be tested prior to installation while the equipment is temporarily insulated from ground.
10.3.4.2 The leakage current flowing through the ground conductor of the power supply connection to ground of permanently wired appliances installed in general or critical care areas shall not exceed 10.0 mA(ac or dc) with all grounds lifted.
10.3.5 Touch Current — Portable Equipment.
10.3.5.1* Touch Current Limits. The touch current for cord connected equipment shall not exceed 100 μA with the ground wire intact (if a ground wire is provided) with normal polarity and shall not exceed 500 μA with the ground wire disconnected.
Testing & Considerations
48

## Slide 47: NFPA99
10.3.5.4 Touch Leakage Test Procedure. Measurements shall be made using the circuit, as illustrated in Figure 10.3.5.4, with the appliance ground broken in two modes of appliance operation as follows:
  (1) Power plug connected normally with the appliance on
  (2) Power plug connected normally with the appliance off (if equipped with an on/off switch)
10.3.5.4.2 Test shall be made with Switch A in Figure 10.3.5.4 closed.
10.3.6* Lead Leakage Current Tests and Limits — Portable
Equipment.
10.3.6.1 The leakage current between all patient leads connected together and ground shall be measured with the power plug connected normally and the device on.
10.3.6.2 An acceptable test configuration shall be as illustrated in Figure 10.3.5.4.
10.3.6.3 The leakage current shall not exceed 100 μA for ground wire closed and 500 μA ac for ground wire open.
Testing & Considerations
49

## Slide 48: Quick view table
Touch current
Testing & Considerations
50
10.2.6* Touch Current — Portable Equipment. The touch current for cord connected
equipment shall not exceed 500 μA with normal polarity and the ground wire disconnected (if
a ground wire is provided).

## Slide 49: Lead Leakage
Lead leakage testing is conducted by connecting all patient leads together and measuring the current from that point to ground
  Ground closed and power on
  Ground open and power on
Portable Equipment
Testing & Considerations
51

## Slide 50: Quick view table
Patient leads
Testing & Considerations
52

## Slide 51: Quick view table
Patient leads
Testing & Considerations
53

## Slide 52: Quick view table
Patient isolation
Testing & Considerations
54

## Slide 53: Class II NFPA 99
NFPA 99 no longer contains instructions for testing equipment with no exposed conductive surfaces
Ungrounded, plastic-encased devices intended for use in the patient vicinity
Should be visually inspected for damage and to verify that it is labeled as double insulated
Testing & Considerations
55

## Slide 54: Non-patient Care-related equipment
Non-patient Care-related equipment, facility or patient owned, shall be visually inspected by staff or “other personnel”
This is a new initiative, basically a common-sense approach. If it appears not to be safe or inoperable it shall be removed or reported
NFPA 99
Testing & Considerations
56

## Slide 55: Visual Inspection
Housing – Enclosure
  Including decontamination
Electrical Connections
  Secondary ground connections
Appropriate fuse rating
Safety markings and labelling
Integrity of mechanical parts
Latest operational hardware and software
Testing & Considerations
57

## Slide 56: Secondary ground
1 Ω Alternative Earth Path via data lead to PC
1 kΩ
MD
1 Ω
IL
2nd Earth
100 μA
99.9 μA
0.1 μA
58

## Slide 57: Power cord
Look for broken or bent pins
Look for discolored pins – may indicate loose connections
Look for physical damage to the strain relief
Inspecting the power cord
Testing & Considerations
59

## Slide 58: Power cord
US
Hot – Black
Neutral - White
Ground - Green
International
Hot – Brown
Neutral – Blue
Ground – Green or Green with yellow strip
IPS
Hot – Brown
Neutral – Orange
Ground - Green
Color codes
Testing & Considerations
60

## Slide 59: Test Process
Manufacturer’s recommendations
NFPA 2012 PAGE 383 – Para 10.5.2.1.2
  Removed the previously required 12 and 6 month testing intervals. Now, testing is required before being put into service for the first time and following any repair or modification that might compromise electrical safety.
Testing & Considerations
61

## Slide 60: DATA
Paper based or Electronic?
Structured
Protocols
Comparing data
Trends
Scheduling
Traceability
Testing & Considerations
62

## Slide 61
Summary & Closing Comments
63
Summary & Closing Comments

## Slide 62: Recommended Reading NFPA 99 2015
Summary & Closing Comments
64
https://catalog.nfpa.org/NFPA-99-Health-Care-Facilities-Code-P1134.aspx?order_src=C900&gclid=EAIaIQobChMIuZSh5pb02gIV2JkbCh0bQw9OEAAYASAAEgLUYvD_BwE

## Slide 63: Recommended Reading
Summary & Closing Comments
65
http://www.aami.org/productspublications/ProductDetail.aspx?ItemNumber=2404
Sponsored by Rigel Medical

## Slide 64: Recommended Reading Guide to 62353
Request your free guide. Ask for a hard copy or download.
  813-886-2775
  sales@seaward-groupusa.com
Summary & Closing Comments
66

## Slide 65: Recommended reading
Summary & Closing Comments
67

## Slide 66
Q&A
Summary & Closing Comments
63
68
Brittany Schmidke
863-370-8774
brittanys@seaward-groupusa.com
Amber Spangenburg
813-335-3345
ambers@seaward-groupusa.com
Jack Barrett
860-729-9972
jackb@seaward-groupua.com