## Slide 1
Your AEM Program: Patient Safety, Standards Compliance, and Less PM!
Matt Baretich
Baretich Engineering, Inc. Fort Collins, Colorado

## Slide 2
What is an AEM program?
CMS: Alternate Equipment Management
TJC: Alternative Equipment Maintenance
CMS premise: Manufacturer knows best
CMS & TJC communication: Inconsistent

## Slide 5
Why implement an AEM program?
Reduce PM cost, consistent with the following criteria
Equipment safety equivalent to manufacturer recommendations
Compliance with CMS and TJC requirements

## Slide 6
Definitions

## Slide 7
Definitions
PM = Planned Maintenance
PM Activity = What we do
PM Frequency = When we do it

## Slide 8
Definitions
PM (Planned Maintenance) includes:
Device Restoration PM
Safety Verification PM

## Slide 9
Definitions
Device Restoration PM
Intended to reduce the likelihood of failure
Genuine preventive maintenance

## Slide 10
Safety Verification PM examples
Personal: blood pressure check, colonoscopy
Personal vehicle: check oil level, check tire pressure
Diesel motor-generator: test oil, monthly load test
Medical equipment ...

## Slide 11
Definitions
PM Procedure = A set of PM activities and PM frequencies
PMOEM = PM that follows the procedure recommended by the manufacturer
PMAEM = PM that follows an AEM procedure

## Slide 12
PM procedure for the Acme 123 (Step 1)
PMOEM – Replace the beaucoup-bucks widget every 24 months
PMAEM – Take a look at the widget every 24 months and replace it only if it’s bad

## Slide 13
PM procedure for the Acme 123 (Step 2)
PMOEM – Verify the flow rate every 12 months
PMAEM – Verify the flow rate every 18 months

## Slide 14
PM procedure for the Acme 123 (Step 3)
PMOEM – Give 3 cheers to the manufacturer every 12 months
PMAEM – Don’t bother

## Slide 15
The Rules

## Slide 16
Table 1. HAP EC.02.04.01 EP4
The hospital identifies the activities and associated frequencies, in writing, for maintaining, inspecting, and testing all medical equipment on the inventory. These activities and associated frequencies are in accordance with manufacturers’ recommendations or with strategies of an AEM program.

## Slide 17
Table 1. HAP EC.02.04.01 EP4
Note 1: The strategies of an AEM program must not reduce the safety of equipment and must be based on accepted standards of practice, such as the ANSI/AAMI handbook ANSI/AAMI EQ56:2013, Recommended practice for a medical equipment management program.

## Slide 18
Table 1. HAP EC.02.04.01 EP4
Note 2: Medical equipment with activities and associated frequencies in accordance with manufacturers’ recommendations must have a 100% completion rate.

## Slide 19
Table 1. HAP EC.02.04.01 EP4
Note 3: Scheduled maintenance activities for both high-risk and non-high-risk medical equipment in an alternative equipment maintenance (AEM) program inventory must have a 100% completion rate. AEM frequency is determined by the hospital’s AEM program.

## Slide 20
On-Schedule PM Completion of Scheduled Maintenance |  | 
 | PMOEM | PMAEM
High-Risk | 100% | 100%
Non-High-Risk | 100% | 100%
Requirements for On-Schedule PM Completion

## Slide 21
Sidebar: On-Schedule PM Completion

## Slide 23
Getting to 100%
  Document legitimate unavailability: cannot locate, in use, etc.
  Only medical equipment: diagnosis, treatment, monitoring
  AEM program to take some off PM schedule
  Wide completion window*

## Slide 24
PM completion window TJC standardization MAYBE
  Scheduled maintenance interval greater than 6 months: scheduled month ± 30 days.
  Scheduled maintenance interval of 3 to 6 months: scheduled month ± 20 days
  Scheduled maintenance interval of 1 or 2 months: scheduled month.

## Slide 25
Taboo List

## Slide 26
Table 2. HAP EC.02.04.01 EP5
The hospital’s activities and frequencies for inspecting, testing, and maintaining the following items must be in accordance with manufacturers’ recommendations:

## Slide 27
Table 2. HAP EC.02.04.01 EP5
Equipment subject to federal or state law or Medicare conditions of participation in which inspecting, testing, and maintaining must be in accordance with the manufacturers’ recommendations, or otherwise establishes more stringent maintenance requirements

## Slide 28
Table 2. HAP EC.02.04.01 EP5
Medical laser devices

## Slide 29
Table 2. HAP EC.02.04.01 EP5
Imaging and radiologic equipment (whether used for diagnostic or therapeutic purposes)

## Slide 30
Table 2. HAP EC.02.04.01 EP5
New medical equipment with insufficient maintenance history to support the use of AEM strategies

## Slide 31
How to decide what to include

## Slide 32
Table 3. HAP EC.02.04.01 EP6
A qualified individual(s) uses written criteria to support the determination whether it is safe to permit medical equipment to be maintained in an alternate manner that includes the following:

## Slide 33
Table 3. HAP EC.02.04.01 EP6
How the equipment is used, including the seriousness and prevalence of harm during normal use
Likely consequences of equipment failure or malfunction, including seriousness of and prevalence of harm

## Slide 34
Sidebar: How to talk about risk

## Slide 35
Risk = Severity x Probability

## Slide 36
AEM-related risk assessments
Overall risk level of a device type: Is this infusion pump a high-risk device?
Risk levels for failure modes (FMEA): What is the risk of battery failure for this infusion pump?
Risk for PMAEM procedures: Is it safe to do flow-rate testing less often?

## Slide 37
Process

## Slide 38
Step 1: Select equipment for your AEM program
  Can you save PM time?
  Is it on your taboo list?
  What is the level of risk?

## Slide 39
Step 2: Review the PMOEM
  Are there PM activities that can be safely modified or eliminated?
  Are there PM frequencies that can be safely extended?
  Are there safe alternative PM strategies?

## Slide 40
Step 3: Implement your PMAEM
  Spell out the modified PM procedure
  Get the modified procedure approved
  Update your CMMS
  Communicate with front-line HTM personnel

## Slide 41
Step 4: Monitor equipment safety
  Routinely monitor your equipment safety metric
  Periodically revisit your PMAEM decisions
  Adjust the AEM program as needed

## Slide 42
Sidebar: How to measure “equipment safety” for your PM program

## Slide 43
Useful metrics
  Patient incidents
  Downtime: Hours per device per year
  Availability: Hours per device per year ÷ required service hours

## Slide 44
Useful metrics
  Failure rate: Failures per device per year
  Mean time between failures (MTBF)

## Slide 45
Nerd Tip: Statistical Process Control (SPC)

## Slide 46
PM-related Failures
  Obvious failures: Recognizable by users; potentially preventable
  Example: Failure of infusion pump battery
  Example: Ventilator component wear-out
  Hidden failures: Not user-recognizable; patient care impact
  Example: Low defibrillator output
  Example: Occlusion pressure sensor failure

## Slide 47
Mean Time Between PM-Related Failures

## Slide 48
Quality assurance for the PM program
  Monitor MTBFPM over time by device type
  Focus on high-risk device types
  Consider using SPC tools to identify problems
  Respond to identified problems
  Adjust AEM procedures as needed
  Periodically revisit AEM decision-making

## Slide 49
Matthew F. Baretich, P.E., Ph.D. Baretich Engineering, Inc. P.O. Box 469 Fort Collins CO 80522-0469 970-493-8239 mfb@baretich.com www.baretich.com