## Slide 1
Presented by
John DiPasquale
Troubleshooting Today’s
X-ray Systems
ISO 9001:2015 CERTIFIED

## Slide 2: A little about me…
Got my start in electronics while in the USAF
  Aircraft electrician – 4 yrs.
  BMET – 18 yrs.
    Technician, Manager, Instructor and Regional Manager
After military service
  8 years in various imaging and biomedical capacities including asset management for both in-house and 3rd party contract positions
  Left industry to teach HS Mathematics in the Connecticut Vocational Technical HS System
  Came back to industry in 2006
    1st OEM experience with Hologic
    Technical trainer on both analog / digital imaging systems and support systems for their mammography lines
  Moved to Wisconsin in early 2013 and began career with TP in December of that year.
Education & Credentials
  Masters degree in Education
  Bachelors degree in Electronics Engineering and Business Administration
  AAS in Instructor and Biomedical Equipment Technologies
  ICC CBET since 1994
  USAF Master Instructor
ISO 9001:2015 CERTIFIED

## Slide 3: What we are going to cover…
X-ray System’s Subsystems
Troubleshooting

## Slide 4: X-ray System’s Subsystems…
Generator Systems
Imaging Systems

## Slide 5: Generator Systems
The Basics…

## Slide 6: Generator Systems(2)…
Required Circuits
Power
  Line to transformer
  230VAC to 480VAC
  Single or three phase
PFT / SFT
  Develops mA
    Direct & indirectly heated cathodes
    Negative side of tube
    Split regulation
      Filament preheat & boost
      mA amount during x-ray
  Older days
    Step down transformer, 200:1
    AC voltages & line frequency
  Currently
    Rectifier & frequency controlled inverter circuits on primary side
    Rectified secondary
PHT / SHT
  Controls length of x-ray
  Develops kV
    Split Secondary
    Phase shifting
    Regulated
  Older days
    Step up transformer, 1:500 ratio
    AC voltages & line frequency
  Currently
    Rectifier & frequency controlled inverter circuits on primary side
    Rectified secondary

## Slide 7: Generator Systems (3)…
Specialty Circuits
Rotor circuits
  Phase shift caused by capacitor start motor circuit
  Speeds
    Low (≤ 4500 RPM)
    High (≥ 10K RPM)
  Start circuits
    AC
    Pulsed DC
  Brake circuits
    DC
    Temp electro magnet
Interlock Circuits
  Monitors all conditions set by law or the operator
    Techniques
    Environmental
  Prevents exposure start
    Unless all conditions are satisfied
    Types:
      Relay logic
      And gate logic
      Microprocessor controlled today
Timer circuits
  Controls the length of the x-ray
  Radiographic (high dose x-rays)
    2 timers
      Simultaneous start
      Primary (5 seconds maximum – radiographic mode)
      Backup (energizes after ~6 seconds)
  Fluoroscopic (low dose x-rays)
    5 minute & resettable

## Slide 8: Generator Systems (4)…
Generator Types
Single Phase
  Half wave
  Full wave
Three Phase
  6 pulse
  12 pulse
High Frequency
  Most common today
  Pulsating DC
Constant Potential
  Generally found in portable systems if used
Voltage Waveforms

## Slide 9: Generators – Single Phase, Full Wave

## Slide 10: Generators – Three Phase, 6 pulse

## Slide 11: Generators – Three Phase, 12 pulse

## Slide 12: Generators – High Frequency

## Slide 13: Imaging Systems…
Radiographic Imaging Systems – Types & Components
Film Systems Components
  3D Top & Tube Stands
    Tube & Collimator positioning
    X, Y, & Z axis movements
    α & β rotations
  Table / Wall Stand
    Phototimers
    Bucky/Grids
  Film
    Direct
    Screen type - Single/double emulsion
  Cassettes
    Paper
    Screen
  Processor
    Manual
    Automatic
  Viewing
    Light Boxes
    Hot lights
  Film Storage Rooms
    Size
    Temp & Humidity controlled
Digital Systems Components
  3D Top & Tube Stands
    Tube & Collimator positioning
    X, Y, & Z axis movements
    α & β rotations
  Table / Wall Stand
    Phototimers
    Bucky/Grids
  Detectors
    Direct
    Scintillation layer type
    Flash circuits
    Types
      Fixed
      Wireless
  Computers
    Processing software
    Image Storage
  Monitors
    Technologist
    Physician
  PACS
    Storage & Retrieval
CR Systems Components
  3D Top & Tube Stands
    Tube & Collimator positioning
    X, Y, & Z axis movements
    α & β rotations
  Table / Wall Stand
    Phototimers
    Bucky/Grids
  Cassettes
    Intensifying screen
    Plate
  Reader
    Laser
    PMT / camera
    Light - flashing
  Computer
    Processing software
    Image Storage
  Monitors
    Technologist
    Physician
  PACS
    Storage & Retrieval

## Slide 14: Imaging Systems (2)…
Radiographic Imaging
Key Terminology
  SID – Source to Image Distance
  SOD – Source to Object Distance
  OID – Object to Image Distance
    Magnification of Image
  HVL – Half Value Layer
    How much aluminum is needed to cut dose in half
    Used to measure changes in inherent filtration
  Focal Spot
    Actual
      Physical area bombarded by electrons
    Effective
      Geometric derived value of central ray used in determining size
      Filaments are rated by this.

## Slide 15: Imaging Systems (3)…
Fluoroscopic Imaging Systems – Types & Components
Film Systems Components
  Spot film device
    Holds I.I. device
    Phototimer
    Film
    Electro-mechanical device
  Image Intensifier Tube
    Converts x-rays to light energy & brightens image
  Light Sensor
    Part of auto kV circuit
    PMT
  Mirror Optics
    Directs image to different viewing /recording devices
  Pick up devices
    Camera tubes
  Monitors
Digital Systems Components
  Spot Film Devices
    Houses detector
    Electronic image capture
  Detectors
    Direct
    Scintillation layer type
    Flash circuits
    Types
      Fixed
      Portable
  Light Sensor
    Part of auto kV circuit
    PDA
  Pick up devices
    CCD cameras
  Computers
    Processing software
    Image Storage
  Monitors
CR Systems Components
  Spot film device
    Holds I.I. device
    Phototimer
    CR cassette
    Electro-mechanical device
  Image Intensifier Tube
    Converts x-rays to light energy & brightens image
  Light Sensor
    Part of auto kV circuit
    PMT / PDA
  Mirror Optics
    Directs image to different viewing /recording devices
  Pick up devices
    Camera tubes
    CCD cameras
  Monitors

## Slide 16: Imaging Systems (4)…
Fluoroscopic Imaging
Key facts…
  Generally real time exams
  Typically uses an under-table tube and radiation is shooting upward
  Exam dependent
    May require use of contrast media
    x-ray altering media that either positively or negatively affects the image depending on what is used to make something standout
  X-ray tube & Imaging Device mounted on C-arms to maintain SID

## Slide 17: Troubleshooting Today’s X-ray Systems…
General Info & Thoughts
Radiographic
Fluoroscopic
Preventive Items

## Slide 18: Troubleshooting…
General Information & Thoughts
Discreet devices evolution
  Individual components
  Single purpose PCBs
  Multiple purpose PCBs
4 key checks
  Signal in
  Signal out
  Power
  Reference
    Earth – Ground (Protective)
    Circuit
      AC
      DC
    Floating

## Slide 19: Troubleshooting (2)…
General Information & Thoughts (cont.)
Analysis questions to ask yourself
  What was it told to do?
  Did it do it?
Another set of questions
  Proper squiggly in?
  Proper squiggly out?
    Size & shape

## Slide 20: Troubleshooting (3)…
General Information & Thoughts (cont.)
X-ray machines are VERY logical
  These systems must have things happen in a sequential order
  It is a building process
  Figure out what is missing in the process and focus on that
  Types of “logic” used over the years
    Relay
    And/Or Gates
    Microprocessors & programs
      Microprocessor looks for information (feedback)
      Time period responses
Understand the limitations of the test equipment you are using…
  Example: Oscilloscope vs DMM
    Both read voltage
    Both are connected in parallel to circuit
    Use circuit reference NOT ground
    Difference is the time based response
      Oscilloscope
        Voltage at an instant in time
        Time per division
      DMM
        RMS voltage over time
        Sampling rate

## Slide 21: Troubleshooting (4)…
General Information & Thoughts (cont.)
Understand the components PRIMARY function in the system
  Generator
    Produce kV & mA
    Control exposure length
  Control Panel / Operator Console
    Sets techniques
      OGP
      Manual
    Environment
  Table
    Patient height
    May have top adjustments
      Motorized
      Floating
    May or may not tilt
      90/90
      90/(15-20)
    2 key signals for generator
    Houses Bucky/Grid assemblies - movement
Understand the components PRIMARY function in the system (cont.)
  Wall Stand
    May or may not - tilt or be motorized
    Houses Bucky/Grid assemblies
    Similar inputs as table
  Stand Movements (Tube & Detector)
    3D Movements – X, Y, & Z axis
    Position Movements – α & β
    Brakes & Tube positioning detents
  Collimator
    Light field represents x-ray field
    May be manual / automatic
      Automatic – PBL / ACSS associated acronym
      CAN system
    Typically 4 blades w/ 2 motors
    May have IRIS function – 8 blades w/ 4 motors (Fluoro)
  X-ray Tube
    Produces the x-rays
    Houses the rotor

## Slide 22: Troubleshooting (5)…
General Information & Thoughts (cont.)
The unwritten rule of troubleshooting…
  If you don’t know what you are looking for – stay out of the machine!!!
  Justifying what the machine tells you will only lead you down the rosy red path to nowhere & probably give you a headache from beating your head up against the wall…

## Slide 23: Troubleshooting (6)…
Radiographic Systems Suggestions
Get as much information from the end user as possible describing what is wrong with the system or about the problem.
  Do not be afraid to ask questions and get clarification if you aren’t sure what you are being told, especially if the message isn’t clear cut
    Table won’t go down
    The image isn’t clear, blurry, has lines in it
Allow the system to help you identify what may be happening and causing your issue
  Example:
    Generate a table top x-ray
    Generate a table x-ray
    Generate a photo-timed x-ray
  Know the differences on how the system is expected to perform and focus on those areas.
    What was added/removed by the environment change

## Slide 24: Troubleshooting (7)…
Radiographic Systems Suggestions (cont.)
Half split methodology
  Generator vs Imaging System
    Items needed
      Intensifying screen
      kV meter
      mA/mAs meter
    Generate an x-ray
      Use either a 2 or 3 point technique
      Screen glows – x-rays coming out
      kV meter – indicates selected kV
      mA/mAs meter – indicates amount of selected mAs
    If all are good, problem will generally lie in the imaging system, NOT the generator
If in the generator
  Figure out what’s missing
    If kV is missing
      All interlocks satisfied
        2 key external interlock signals
          Cassette present
          Bucky movement
      Are items reset
      Error codes, information available
    If mA/mAs missing
      Filaments open
      Proper filament regulation during standby & boost
      Proper regulation for mA during x-ray
    Are the power supplies proper and working

## Slide 25: Troubleshooting (8)…
Radiographic Systems Suggestions (cont.)
If problem is in the Imaging System, then the type of Imaging System determines what needs to be checked/done
  Film Based
  CR Based
  Detector Based
Film Based
  Film Processor biggest culprit
    Compare sample against daily QC results for processor
      If not within range – processor needs attention
      Generally speaking
        Chemical strength
        Temperature issues
    Optical Density (OD) not proper
  Dark Room
    Safe light
    Chemical fumes
    Storage cabinet
  Cassettes
    Poor contact
    Light leaks

## Slide 26: Troubleshooting (9)…
Radiographic Systems Suggestions (cont.)
CR Based
  S-curve value
    Must be ≥ minimum value for reader to work
    Similar in nature to OD value used in film but generally higher
  Reader
    Not clearing plate
      Flash deteriorating
      Ghosted image (double exposure)
    Defective laser
    Defective pickup device
  Computer
    Data transfer
    Corrupt program
    Defective computer
  Monitor
    SMPTE adjustment
    Defective display

## Slide 27: Troubleshooting (10)…
Radiographic Systems Suggestions (cont.)
Detector based
  Detector files
    Correct & current files
    Latest mapping/calibration files
  Detector
    Power supply
    Flash deteriorated or not working
    Defective pixels
    Bad cables (fiber optic)
    Wireless
      Defective battery
      Not paired
  Computer
    Data transfer
    Corrupt program
    Defective computer or PCB
    Files corrupt, missing or incorrect
  Monitor
    SMPTE adjustment
    Defective display

## Slide 28: Troubleshooting (11)…
Fluoroscopic Systems Suggestions
All previous items discussed still apply
  Half split methodology
  Generator info
Two key additional checks to add on the generator side though
  Max R output of the tube ensuring the generator is doing its part
    Ensure you measure it according to manufacturer’s specifications
    Use the probe sensor indicator to set distance
  Use different items of thickness & density in the beam
    Verify the auto kV circuit is working
On the Imaging System side one has to keep in mind this is live imaging – not a single shot
  Image Intensifier Tube
    Biggest factor are the voltages to the acceleration plates
    High voltage
      Requires a HV probe for your meter
      Some systems have representative low voltage test points for this
    Use a mirror to check the phosphor output of the tube
  Digital Detector
    High Voltage to Detector
    Use mirror to check the phosphor output of the detector

## Slide 29: Troubleshooting (12)…
Fluoroscopic Systems Suggestions (cont.)
Collimator
  Fluoro typically has an 8 blade collimator effectively creating a octagonal (circular) input to the I.I.
  Smaller the field, the larger the magnification of the item being examined.
    Most systems have 3 different levels of magnification
  Most common problems
    Motors
    Calibration
PMT / PDA
  Typically positioned on the output side of the I.I. / detector
  Used to sense the amount of light coming out of the I.I. / detector
  Part of the auto kV adjustment circuit used to ensure the correct amount of light is being sent to the pickup device.
    Alters penetration of the x-rays by changing kV
    Changes output light back to “normal” value
Pickup devices
  Types
    Video tubes
    CCD cameras
  Takes the light from the I.I./detector and sends it off to the computer for processing as digital data
    Video tubes output degrades over time
    CCD cameras
      Require defect maps for computer processing
      Bad pixels
      Camera Iris not working correctly
Mirror Optics
  Directs the output of the I.I./detector to different devices for the purpose determined by the radiologist.
  Mirror movement
  Light out

## Slide 30: Troubleshooting (13)…
Fluoroscopic Systems Suggestions (cont.)
Spot Film Devices (C-arm)
  Older days
    Electro-mechanical nightmares
      Movement controls for adjusting the c-arm mechanism during the exam
    Positions film for a radiographic exposure based on programming pattern
      1x1, 2x1 (L/H), 4x1, 6x1, 9x1, 12x1, 16x1
    Temporarily shifted operation of the system from fluoroscopy to radiography and then back again when completed
      Records what Doctor saw during an exam
      Higher mAs & dose to patient
      Phototimed x-ray
      Collimator leaves Iris mode goes to rad mode, then back to Iris mode
      Brakes c-arm during rad exposure
Spot Film Devices (C-arm) (cont.)
  Newer devices
    Movement control for positioning the c-arm over the patient during the exam
    Does an “image capture” of the fluoroscopic image that is seen by the doctor
      Multiple screen systems can show both the live fluoro exam and the most recent image capture
      Lower radiation to patient as the fluoro image is what is captured
      Less mechanical devices for FE’s to concern themselves with
      More remote movement controls for operator’s to use creating multiple ways of controlling the movements of the components.
    All images processed by computer on the system

## Slide 31: Troubleshooting (14)…
Fluoroscopic Systems Suggestions (cont.)
Computers
  Input connections
    Fiber Optic Cables used with Imaging
    Keyboards
    Mouse
  Power supplies
  HDDs
    Damaged
    Corrupt files
    Maps & Calibration files
  Video cards
    Image processing
Monitors
  Calibrations
  SMPTE Patterns
  VGA/HDMI outputs
  Color / Monochrome images
    Use separate monitors
    Set color for B&W (monochrome) if desired
  Power supplies
  Proper pixelization
    Doctor vs technologist monitors

## Slide 32: Troubleshooting (15)…
Suggested Preventive Items for the Field Engineer
Cloning computer hard drives
  Suggested programs to use
    Norton Ghost
    CloneZilla
  Frequency
    After initially installed
    Changes to programming
    Any software upgrade
    Annual PM
Backups
  At the conclusion of a PMI
  Things to consider
    Generator
    System calibrations
      3D Top / Floor Stand
      Table
      Collimator
      Wall Stand
    Computer information
When performing service
  Do a backup before you begin
  Perform the work necessary
  Do a backup after the service is done
  Update info left on site
For those of you servicing multiple sites
  Sharing information between all engineers responsible for the site
  Do NOT solely rely on the internet for access to this information

## Slide 33
Contact Information
Technical Prospects
1000 County Road CB
Appleton, WI 54919
Main #: (920) 757-6583
Toll Free #: (877) 604-6583
Fax #: (920) 757-6591
www.technicalprospects.com
My information:
  John DiPasquale, ext. 8480 jdipasquale@technicalprospects.com
Parts/Sales:
Customer Service Reps, ext. 2101
  parts@technicalprospects.com
Training Information:
  Linda Fuerst, ext. 8202
  training@technicalprospects.com